- RICE UNIVERSITY

- DEVELOPMENT OF AN ION/ELECTRON PLASMA SPECTROMETER

by

- WAYNE R. KEITH

- A THESIS SUBMITTED

IN PARTIAL FULFILLMENT OF THE

REQUIREMENTS FOR THE DEGREE

MASTER OF SCIENCE

- APPROVED, THESIS COMMITTEE

__________________________________

- J. David Winningham

Adjunct Professor, Space Physics and

Astronomy

Institute Scientist, Southwest Research

Institute

_________________________________

- Patricia H. Reiff

Chairman, Space Physics and Astronomy

_________________________________

- Jon C. Weisheit

Adjunct Professor, Space Physics and

Astronomy

Division Director, Los Alamos National

Laboratory

- Houston, Texas

April, 1999

Abstract

Development of an Ion/Electron Plasma Spectrometer

by

Wayne R. Keith

This thesis details the development of the Miniaturized Electrostatic DUal top-hat Spherical Analyzer (MEDUSA) instrument, including calibration, simulation, and data processing software.  MEDUSA combines two-species observations with state-of-the-art miniaturization to achieve a low-cost, compact instrument capable of making the measurements that will be required in the future.  Calibration of the instrument was performed during November 1997 and July 1998.  In order to better understand the characteristics of the instrument, it was also modeled in detail with a modified 2.5 dimensional raytracing code.  The production processing software is designed to archive the data into the standard Instrument Data File Set (IDFS) format of the Southwest Data Display and Analysis System (SDDAS).  MEDUSA has so far proven itself to be a reliable, low-cost instrument capable of making the scientific measurements required for a better understanding of the behavior of the near-Earth particle environment.

Acknowledgements

There have been many people who have influenced the course of my life that has led to this thesis.  As I was growing up, I am glad my parents encouraged me to follow my dreams, even if they did think it was just a phase.  My high school Physics teacher, Stanley Miller, showed me how exciting it can be to try to understand the world around us.  As an undergraduate, Dr. Jimmy McCoy helped me to focus my dreams into a roadmap to where I wanted to go, and the Mathematics and Physics staff at Tarleton got me started on that road.  At Rice, the Space Physics and Astronomy faculty have been superb, both in and out of the classroom.  A special thanks to Drs. Patricia Reiff and Jon Weisheit, who in addition to serving on my committee were of great value during my on-campus days.

The people who have had the most profound impact on this thesis, however, were the employees of Southwest Research Institute in San Antonio, Texas and at the Swedish Institute for Space Physics in Kiruna, Sweden.  At Southwest, I'd like to thank my advisor, Dr. David Winningham, for all his help on my thesis and for getting me into the kinds of projects I wanted to be involved with.  I would also like to thank Dr. Rudy Frahm for all his advice and proofreading, not to mention the occasional well-cooked meal.  Many others at Southwest have helped me along the way, including Dr. Jim Sharber, John Scherrer, and Gene Horton.  At IRF, Dr. Olle Norberg was very helpful with calibration issues.

Special thanks go to my friends and family who have encouraged me along the way.  My Aunt Billye, who inspired me to dream big, and my Uncle Lester, who always made sure I knew what things’ proper names were.  Special thanks go also to my parents and grandparents for knowing how hard to push, and my fiancée Melinda for always being there.  Finally, I give thanks to God, without whom none of my dreams would prosper.

Finally, I'd like to dedicate this thesis to Stanley Miller, mentor and friend, and to Aunt Billye, my biggest fan, who both passed away during its preparation.

Table of Contents

Abstract…………………………………………………………………………….…….ii

Acknowledgements………………………………………………………………….…..iii

Table of Contents………………………………………………………………….……..v

List of Figures…………………………………………………………………….…….vii

Abbreviations………………………………………………………………………….…x

1.  Introduction………………...………………………………………………………….1

1.1.  Scientific Objectives……………………………………………………………...1

1.2.  The Astrid-2 Mission……………………………………………………………..3

1.3.  The Munin Mission…………………………………………………………….…5

1.4.  Background…………………………………………………………………….….8

1.5.  The Top Hat Analyzer………………………………………………………….…9

1.6.  The Particle Sensor………………………………………………………………11

2.  Instrument Design…………………………………………………………………….14

2.1.  Physical and Environmental Constraints………...………………………………14

2.2.  Design Characteristics…………………………………………………………...17

2.3.  Design Overview………………………………………………………………...20

3.  Instrument Development……………………………………………………………...28

3.1.  Engineering Model…………………………………………………………….....28

3.2.  Prototype/Flight Model…………………………………………………………..30

3.3.  Bench Testing..…………………………………………………………………..39

3.4.  Chamber Testing...……………………………………………………………….42

3.5.  Troubleshooting at IRF…………………………………………………………..48

3.6.  Second Flight Model……………………………………………………………..51

4.  Calibration……………………………………………………………………………53

4.1.  Facility…………………………………………………………………………...53

4.2.  Procedure and Results……………………………………………………………55

5.  Raytracing Model…………………………………………………………………….74

5.1.  Model Setup and Parameters…………………………………………………….74

5.2.  Comparison with Calibration Data………………………………………………77

5.3.  Collimator Study Results………………………………………………………...87

6.  Software Development………………………………………………………………114

6.1.  SDDAS Overview…………………………………………………………….114    6.2.  Production and Viewing Software……………………………………………...116

7.  Conclusions………………………………………………………………………….120

7.1.  Scientific Value…………………………………………………………………120

7.2.  Educational Value………………………………………………………………121

References…………………………………………………………………………...…122

WWW References……………………………………………………………………...123

Appendix A:  Distance Between Non-Concentric Circles……………………………...124

Appendix B:  Scaling of Centaur Characteristics………………………………………126

Appendix C:  MEDUSA Data Processing Code (Sample)……………………………..129

Appendix D:  MEDUSA Electron VIDF File..................................................................148

Appendix E:  MEDUSA Electron PIDF File..……………………………………….…153

Table of Figures

1.  Introduction

1.1.  Scientific Objectives

1.2.  The Astrid-2 Mission

Figure 1-1 – The Astrid-2 Satellite

1.3.  The Munin Mission

Figure 1-2 – The Munin Satellite

1.4.  Background

Figure 1-3 – Top Hat Schematic Drawing

1.5.  The Top Hat Analyzer

Figure 1-4 – Top Hat Sectional View

1.6.  The Particle Sensor

2.  Instrument Design

2.1.  Physical and Environmental Constraints

Figure 2-1 – Astrid-2 Satellite Upper Platform

2.2.  Design Characteristics

Figure 2-2 – MEDUSA Sectional Diagram

2.3.  Design Overview

Figure 2-3 – MEDUSA Sensor Body

Figure 2-4 – MEDUSA Amplifier Board

Figure 2-5 – MCP D" Cutout Diagram

Figure 2-6 – MEDUSA Inner Deflection Plate

Figure 2-7 – MEDUSA Alignment Jig (Ion Side)

Figure 2-8 – MEDUSA Connector Cover and Backing Plate

3.  Instrument Development

3.1.  Engineering Model

Figures (3-1 – 3-2) – MEDUSA Engineering Model

3.2.  Prototype/Flight Model

Figures (3-3 – 3-4) – MEDUSA Prototype Flight Model

Figure 3-5 – MEDUSA Electron Amplifier Boards

Figures (3-6 - 3-9) High Voltage Relation Graphs

  - Bench Testing

Figure 3-10 – MEDUSA Bench Setup Diagram

Figure 3-11 – MEDUSA Bench Testing Setup

3.4.  Chamber Testing

Figure 3-12 – Chamber Diagram

Figure 3-13 – Chamber Vacuum Diagram

Figure 3-14 – Ion Source Diagram

3.5.  Troubleshooting at IRF

Figure 3-15 – Modified Lower Spring and Outer Support Ring

3.6.  Second Flight Model

Figure 3-16 – Modified Inner Support Ring

4.  Calibration

4.1.  Facility

Figure 4-1 – Calibration Facility Diagram

Figure 4-2 – MEDUSA Mounted in Vacuum Tank

Figure 4-3 – Argon+ Ion Source and Support Equipment

Figure 4-4 – Electron Source and Pumps

Figure 4-5 – Electron Source and Support Equipment

Figure 4-6 – Calibration Software Screen Shot

4.2.  Procedure and Results

Figure 4-7 – Electron Counting Rate vs. MCP Bias Voltage

Figure 4-8 – Ion Counting Rate vs. MCP Bias Voltage

Figure 4-9 – Electron Counting Rate vs. Beam Intensity

Figure 4-10 – Ion Elevation Acceptance Curve (manual)

Figure 4-11 – Ion Elevation Acceptance Curve (automatic)

Figure 4-12 – Ion Azimuth Scan Integrated over Elevation

Figure 4-13 – Electron Geometric Factor vs. Beam Energy

Figure 4-14 – Ion Geometric Factor vs. Beam Energy

Figure 4-15 – Electron K Factor and Gain vs. Sector (540º)

Figure 4-16 – Ion K Factor and Gain vs. Sector (540º)

5.  Raytracing Model

5.1.  Model Setup and Parameters

Figure 5-1 – Beam Equivalence Diagram

5.2.  Comparison with Calibration Data

Figure 5-2 – Realistic MEDUSA Model

Figure 5-3 – Simple MEDUSA Model

Figure 5-4 – Geometric Factor vs. Input Energy, Simple Model

Figure 5-5 – Geometric Factor vs. Input Energy, Realistic Model

Figure 5-6 – Geometric Factor vs. Elevation Angle, Simple Model

Figure 5-7 – Geometric Factor vs. Elevation Angle, Realistic Model

Figure 5-8 – Initial Spacing – Incorrect Inner Radius

Figure 5-9 – Spacing After Addition of Shims

5.3.  Collimator Study Results

Figure 5-10 – Distribution of Successful Input Locations, Realistic Model

Figure 5-11 – Single Mask Model

Figure 5-12 – Distribution of Successful Input Locations

Figure 5-13 – Geometric Factor vs. Input Energy, Single Mask

Figure 5-14 – Geometric Factor vs. Elevation Angle, Single Mask

Figure 5-15 – Dual Mask Model

Figure 5-16 – Distribution of Successful Input Locations

Figure 5-17 – Geometric Factor vs. Input Energy, Dual Mask

Figure 5-18 – Geometric Factor vs. Elevation Angle, Dual Mask

Figure 5-19 – Diagram of Output Mask Model

Figure 5-20 – Geometric Factor vs. Input Energy, Output Mask

Figure 5-21 – Geometric Factor vs. Elevation Angle, Output Mask

Figure 5-22 – Flat Collimator Model

Figure 5-23 – Distribution of Successful Input Locations

Figure 5-24 – Geometric Factor vs. Input Energy, Flat Collimator

Figure 5-25 – Geometric Factor vs. Elevation Angle, Flat Collimator

Figure 5-26 – Flat Collimator Entrance Mask Model

Figure 5-27 – Distribution of Successful Input Locations

Figure 5-28 – Geometric Factor vs. Input Energy, 0º Entrance Mask

Figure 5-29 – Geometric Factor vs. Elevation Angle, 0º Entrance Mask

Figure 5-30 – Distribution of Successful Input Locations

Figure 5-31 – Geometric Factor vs. Input Energy, 0º Separatrix

Figure 5-32 – Geometric Factor vs. Elevation Angle, 0º Separatrix

6.  Software Development

6.1.  SDDAS Overview

6.2.  Production and Viewing Software

Figure 6-1 – Data Production Flowchart

Figure 6-2 – Select Mode Data in IDFS Format

7.  Conclusions

7.1.  Scientific Value

7.2.  Educational Value

Appendix A

Figure A-1 – Diagram of Non-Concentric Circles

Appendix B

Figure B-1 – Side View of Top Hat Plates

Figure B-2 – Top View of Aperture Opening

List of Abbreviations

AC		August, 1998 Calibration

ACS		Attitude Control System

ASC		Advanced Stellar Compass

CCD		Charge Coupled Device

CEM		Channel Electron Multiplier

DoD		Department of Defense

DPU		Data Processing Unit

EM		Engineering Model

ESA		European Space Agency

EsA		Electrostatic Analyzer

FM		Flight Model

FWHM	Full Width at Half Maximum

GUI		Graphical User Interface

IDFS		Instrument Data File Set

IRF		Swedish Institute for Space Physics (Institutet för Rymdfysik)

MCP		Micro Channel Plate

MEDUSA	Miniaturized Electrostatic DUal-top-hat Spherical Analyzer

MSP		Micro Sphere Plate

NASA		National Aeronautics and Space Administration

NC		November, 1997 Calibration

OA		Orbit Attitude

PIA 		Photometers for Imaging the Aurora

PIDF		Plot Information and Description File

SDDAS	Southwest Data Display and Analysis System

RM		Realistic Model

SM		Scaled Model

SwRI		Southwest Research Institute

TFAS		Top-hat For All Species

VIDF		Virtual Instrument Description File